Request Title:  Extended Exposure ALD of Platinum in Savannah Systems

Requestor: 	Donald Gardner					Date:  4/25/2019

Non-SNF staff, please complete the following:

Badger ID:	don.gardner

Company name/PI: SLAC / Anne Sakdinawat

Send completed form to: snf-promcommittee@lists.stanford.edu

- The chemical or material.

  - Precursor: Trimethyl(methylcyclopentadienyl)platinum(IV).

Same precursor currently in use in Fiji systems.

MSDS available on SNF Wiki.

  - Oxygen gas: Need to add to Savannah systems.

- Vendor/manufacturer info:

  - SNF already purchases this chemical from Strem.

- Reason for request:

- Filling high-aspect ratio structures using atomic layer deposition requires long exposure times and high precursor partial pressures.  This results in the consumption of large quantities of precursors as well as usage of the ALD tool for long periods of time. The Savannah system has the following advantages over the Fiji ALD systems for extended exposure ALD:

  - The Savannah chamber has a significantly smaller chamber. Fiji has a chamber volume of 19.1 liters versus 0.4 liters for Savannah. The ideal gas law predicts that the factor of 48 reduction in volume results in the precursor partial pressure increasing by a factor of 48. The product of the pressure P of a gas and the volume V is a constant for an ideal gas.

  - Thermal ALD is the best choice for extended exposure ALD because plasmas do not penetrate well into high-aspect-ratio features and would typically not be used.  To complicate matters, the Fiji ALD systems has a plasma generator that requires a constant flow of argon gas to keep it from being coated. This argon gas flow causes the pressure to increase when a stop flow valve is closed potentially interfering with the deposition process as well as causing concerns for overpressure conditions.

  - The amount of Pt precursor needed would be significantly reduced since the pulse is so much more concentrated. This reduces the time and costs associated with the process.  Note that the Pt precursor is an expensive precursor.

  - The Fiji systems have a high-surface-area honeycomb structure under the substrate holder that getters the residual precursor before it exits into the exhaust to reduce the residual precursor in the exhaust. This structure appears to also getter the precursor during extended exposures in the Fiji systems (confirmed by Cambridge Nanotech). In the Savannah system, the getter is after the stop valve so the precursor during the extended exposure is not gettered until after the deposition time has completed.

- The Fiji ALD is not designed for extended exposure processes whereas the Savannah ALD is.  We have fond that the highest aspect ratios that the Fiji system can fill to be limiting using the ALD Pt extended exposure recipes and that longer exposures actually results in inferior sidewall coverage, contrary to what is predicted by theory and the literature.  In addition, Cambridge Nanotech has stated that “the Savannah is more suited to Expo mode processes.”

- Process Flow:

- The atomic layer deposition will be carried out in either or both of the two Savannah tools, ALD tools that are classified as gold-contaminated.  The Pt is used to coat the high-surface area structures. The surface is first coated with a nucleation layer such as Al2O3. To deposit the Pt, the Pt precursor is first pulsed into the chamber with the stop valve closed for a set period of time such as 10 seconds to 2 minutes, then the stop valve is opened and a purge gas is flowed to remove the precursor.  The second half of the cycle is to introduce oxygen for a similar amount of time and then purge the oxygen.

- The process temperature for the standard ALD Pt thermal process (on the Fiji systems) is 270C. It is desirable to use the same process temperature such that the films from Fiji and Savannah have similar morphologies and characteristics.  The Savannah system is listed on the SNF wiki as capable of temperatures up to 250C, but per Cambridge Nanotech, the tool is capable of going to 300C as long as it has Karlex O-rings or equivalent.  If it uses Viton, then these need to be changed. Another option is to lower the temperature of the outer heaters that are close to the O-rings, but this can result in nonuniformities and adds complexity to the process. Cambridge Nanotech suggested reducing the temperature of the outer heaters for processes that require temperatures above 300C (e.g. 325C) which is not necessary for this ALD Pt process.

- Amount and form.

  - The precursor is already used in SNF. It is purchased in 10 g bottles.

  - Oxygen is already in use on the Fiji ALD systems. It would need to be added to the Savannah systems.

- Storage:

  - The precursor is already stored in SNF.

  - Oxygen is already stored in SNF.

- Disposal:

  - The excess precursor is already dealt with by SNF staff.

To be completed by the PROM COMMITTEE:

Special handling requirements:

On-going approval or staff interaction requirements:

Approved By/Date: